[image: image1.jpg]



CUMnNi resistor/ CT– precision current sensor  Type series EM-KS


[image: image2.png]EasyMete




PRELIMINARY
CUMnNi resistor/ CT– precision current sensor  Type series EM-KS


      [image: image3.png]EasyMete







Applications:


- sensor for power-metering

- Current sensing in 
  automotive applications

- DC resistant current sensor

- free of potential current 
  measurement 

Technical data: of KS200 / KS400

	Main Shunt: resistance 
	0,4 mOhm  for KS400 / 0,2 mOhm for KS200

	Parallel resistor: resistance
	68 mOhm

	CT: Number of windings
	1350

	CT: Typical burden; typical phase shift
	72 Ohm; 0,8 degree at 50Hz

	Tolerance
	+/- 10%

	Temperature coefficient
	< 50 ppm/K (20°C to 60°C)

	Applicable temperature range
	-55 to +125°C

	Load 
	3 Watt

	Max. pulse power
	100 Watt / 0,1 s

	Permanent current (DC or AC)
	60A (KS400) / 100A (KS200)

	Internal heat resistance
	< 10 K/W

	Inductance
	< 1 nH

	Average lin. thermal expansion coefficient
	18 x 10-6/K

	Long term stability (nominal load)
	Deviation < 0.4% after 1000h at TK = 80°C

	Assembly 
	Soldering or welding on copper

	Weight
	approx. 25g


Circuit of the KS200 / KS400


Dimensions for KS200 /KS400:


for easy metering





7





2,5





11





20





32





6,3





11





10,5





10,5





25








IPrim








parallel Resistor





Shunt





Burden





vBu














12,5










CUMnNi resistor/ CT 
precision current sensor Type series EM-KS
EM-KS2_10.11.03_vers. 1.0
page 1


CUMnNi resistor/ CT 
precision current sensor Type series EM-KS
EM-KS2_10.11.03_vers. 1.0
page 2

