[image: image1.png]A\ NEMOTO & CO,, LD,




Electro-Chemical Sensor  Sensor   NAP－505R
NAP-505 gas sensor is our standard type of electrochemical CO gas sensor for a CO detector for domestic use. Detection principle is quite the same as conventional one, but quite cheaper than conventional ones due to unique structure and renovated electrodes. NAP-505 has already been approved by UL. (UL2034 - Residential CO detectors) As there are various types of NAP-505 as shown below, please contact us for detailed information 

Characteristics
	Detection range
	CO 0 ~ 1,000ppm
	Zero point drift
	Below 10ppm(-20~50oC)

	Output current
	40 ± 10nA/ppmCO
	Temp. and humidity
	-20~50oC、15-90%RH

	Repeatability
	Less than ±2%
	Span drift
	Less than 5%/year

	Response time
	T90 : less than 30sec.
	Lifetime
	Over 5 years


	[image: image2.png]ol






	Fig. Gas sensitivity of NAP-505

	[image: image3.png]4
=3
E]
= w
JEsr=
x
i
£ o) =
o 1 F—Jae-Bitave, Uethan
oo
! 1=

0100 20 W0 40 w0 B 700
‘Gas Concentration (pom)




	[image: image4.jpg]





